## Slide 1: ISO 13485:2016 Norm for regulatory compliance
Alice Ravizza

## Slide 2: Risk subjects
Manufacturer is compelled to evaluate impact of use of the device:
  On patient
  On intended user
  On bystanders
  On general environment
Manufacturer is compelled to evaluate effect of product impact in all life cycle, from manufacturing to disposal

## Slide 3: Norms
Device lifecycle is regulated as per:
  ISO 13485 for Quality Systems
  ISO 14971 for Risk Management
  ISO 14155 and various guidelines for Clinical Investigations
Each product category is then regulated by technical norms
  For electro medical devices
  For sterile devices
  For devices in contact with the body
  Multiple harmonised and not harmonised norms for technical regulation

## Slide 4: Quality Management System
Full company management system
  Internal processes
  Processes directed to the customer
  Continuous feedback
For an organization to function effectively, it has to identify and manage numerous linked processes

## Slide 5: Definitions….
Process: an integrated set of activities that uses resources to transform inputs into outputs
System: several processes are interconnected using such input-output relationships.
NOTE: Processes are interconnected to form a system because the output from one process becomes the input for another process.

## Slide 6: Process approach: a management strategy
Managers shall identify and rule:
  processes that make up their organization,
  the interaction between these processes
  the inputs and outputs that glue these processes together
Each process is evaluated by its capacity to reach a set goal

## Slide 7: Typical processes for any company
Quality Management/ Management review
Training
Design and Development
Production Management
Service Provision
Logistic Management
Customer Relationship Management
Product Purchasing/ Supplier qualification

## Slide 8: Manufacturers’ responsibility
Develop the QS and apply regulation as applicable to their specific products and operations
Establish
  requirements for each type or family of devices to ensure that devices are safe and effective,
  methods and procedures to design, produce, distribute, etc. devices that meet the requirements

## Slide 9: General rule for all systems
  The responsibility for
    meeting system and product requirements
    having objective evidence of meeting these requirements
    MAY NOT BE DELEGATED

## Slide 10: Aim of the standard ISO 13485:2016 clause 1.1
Specify requirements on SYSTEM
Facilitate global alignment of appropriate regulatory requirements
Provide an unique, integrated approach
  Evaluation of process risk
  Subsequent management

## Slide 11: Receivers ISO 13485:2016 clause 1.1
Companies that need to implement and maintain a “strong” QMS
  Consistently manage the life cycle of a MD or IVD
  Meet customer requirements
  Meet regulatory requirements
Any company regardless size or type
  Manufacturers of finished devices
  Manufacturers of subassemblies (no regulatory requirements as per MDD or IVD)
  Suppliers of services

## Slide 12: Receivers
organization involved in stages of the life-cycle of a medical device
  including the design and development, production, storage and distribution, installation or servicing of medical devices, and the design, development, and provision of associated activities
suppliers or other external parties providing goods and services
  sterilization services, calibration services, distribution services)

## Slide 13: Mandatory?
  Voluntarily choose to conform to the requirements of this standard
    OR may be required by contract to conform
  As all HARMONISED standards, allows claiming presumption of compliance to Essential Requirements of Medical Device Directive/ IVD Directive

## Slide 14: Role of the company
The organization shall:
  identify its role(s) under appropriate regulatory requirements,
  identify the regulatory requirements that are appropriate for its activities under these roles, and
  incorporate these appropriate regulatory requirements within its quality management system.

## Slide 15: Relationship to other standards
Can be used as an part of an integrated system
  Business 9001 series
  Environment ISO 14000 series
  Health and safety
  Ethics

## Slide 16: Relationship to ISO 9001:2015
Common approach:
  process approach
  risk-based thinking
Different structure
  ISO 13485:2016 still reflects the structure from ISO 9001:2008

## Slide 17
ISO 9001:2015 | ISO 13485:2016
Understanding the context of the organization, its quality management system and processes | Quality management system
Leadership, policy and responsibilities | Management
Processes for planning and consideration of risks and opportunities | Resource management
Processes for support, including resources, people and information | Product realization
Operational processes related to customers and products and services | 
Processes for performance evaluation | Measurement, analysis and improvement
Processes for improvement | 

## Slide 18: Risk based approach
“One size fits all” or “worst case” approach is not feasible
  Technically
  economically
“Sizing” the system on the company
  objectives,
  products / services
  Processes
  size and organizational structure
  appropriate regulatory requirements

## Slide 19: Limitations to applicability
Some product limitations (clause 6, 7 or 8)
  Not sterile (waive 6.4.2 and 7.5.1.3)
  Not cleaned (waive 7.5.1.2.1)
  Not installed (waive 7.5.1.2.2)
  Not serviced (waive 7.5.1.2.3)
  Not active implantable (waive 8.2.4.2)
Exclusion of design activities

## Slide 20: 4 Quality management system
QMS defined and maintained
  According to ISO standard
  According to regulatory requirements
Documentation of role of organization
  manufacturer, authorized representative, importer or distributor.

## Slide 21: System planning clause 4.1.2
Determine processes
Control them via a risk based approach
Determine process flow/interactions

## Slide 22: System functioning clause 4.1.3 and 5.4.2
Criteria and methods to ensure process effectiveness
Resources
Achieve planned results (including quality objectives)
Monitor, measure
Record (written)

## Slide 23: System change control clause 4.1.4
Impact on QMS
Impact on products
Control:
  Non- distruption of QMS (see also clause 5.4.2)
  Compliance to regulatory requirements

## Slide 24: Control of outsourced processes Clause 4.1.5
The organization shall retain responsibility of conformity
Monitor and control
Written quality agreements NEW

## Slide 25: SW validation clause 4.1.6
SW used in the QMS NEW
  First use
  changes
Risk based approach to SW validation
NOTE: written records required
NOTE: CEI 62304

## Slide 26: 4.2 Documentation requirements
Quality policy
Quality objectives
Quality manual
Procedures and records
Documents for planning, operation and control (work instructions)
Others?

## Slide 27: 4.2.2 Quality manual (unvaried)
a) the scope of the QMS,
  justification for any exclusion or non-application;
b) the documented procedures;
  outline the structure of the documentation
c) interaction between the processes

## Slide 28: 4.2.3 Medical device file NEW
For each medical device type or medical device family
Documents generated to demonstrate conformity with
  requirement of ISO
  compliance with applicable regulatory requirements

## Slide 29: Contents of file
a) general description of the medical device, intended use/purpose, and labelling, including any instructions for use;
b) specifications for product;
c) specifications or procedures for manufacturing, packaging, storage, handling and distribution;
d) procedures for measuring and monitoring;
e) as appropriate, requirements for installation;
f) as appropriate, procedures for servicing.

## Slide 30: IVD regulation proposal ANNEX II TECHNICAL DOCUMENTATION
DEVICE DESCRIPTION AND SPECIFICATION, INCLUDING VARIANTS AND ACCESSORIES
INFORMATION SUPPLIED BY THE MANUFACTURER
DESIGN AND MANUFACTURING INFORMATION
GENERAL SAFETY AND PERFORMANCE REQUIREMENTS
RISK/BENEFIT ANALYSIS AND RISK MANAGEMENT
PRODUCT VERIFICATION AND VALIDATION

## Slide 31: Written SOP for document control 4.2.4 Control of documents
a) review and approve documents for adequacy prior to issue;
b) review, update as necessary and re-approve documents;
  Review by the same authority as issue
c) ensure that the current revision status of and changes to documents are identified;
d) ensure that relevant versions of applicable documents are available at points of use;
e) ensure that documents remain legible and readily identifiable;

## Slide 32: Written SOP for document control 4.2.4 Control of documents
f) ensure that documents of external origin, determined by the organization to be necessary for the planning and operation of the quality management system, are identified and their distribution controlled;
g) prevent deterioration or loss of documents;
h) prevent the unintended use of obsolete documents and apply suitable identification to them
controls needed for the identification, storage, security and integrity, retrieval, retention time and disposition of records

## Slide 33: 4.2.5 Control of records (unvaried)
provide evidence of
  conformity to requirements
  the effective operation of the QMS
shall remain legible, readily identifiable and retrievable.
Changes to a record shall remain identifiable
  NOTE: Good documentation practices

## Slide 34: Filing of documents and records
Filing for a predefined period
  Procedures: At least same of medical device lifetime
  BUT
  At least more than the resulting records
  Records: at least 2 years from release
BUT filing time also in compliance to regulatory requirements

## Slide 35: 5.1 Management commitment and 5.2 customer focus
Communication
  Quality policy
Quality objectives
  Management reviews
Availability of resources
Customer focus
Regulatory requirements focus

## Slide 36: 5.3 Quality policy
a) is applicable to the purpose of the organization;
b) includes a commitment to comply with requirements and to maintain the effectiveness of the quality management system;
c) provides a framework for establishing and reviewing quality objectives;
d) is communicated and understood within the organization;
e) is reviewed for continuing suitability.

## Slide 37: 5.4.1 Quality objectives
Quality objectives
  Efficiency, efficacy
  Regulatory requirements
  Product requirements (example: expected AqLs)
Measurable
Consistent with the quality policy

## Slide 38: 5.5.1 Responsibility and authority 5.5.2 Management representative
Define and communicate responsibilities
  Organizational chart
  Job profiles
Personnel affecting quality
  ensure the independence and authority
Management representative
  Responsible of QMs
  Reports to top management
  Promotes the QMS in the organization

## Slide 39: 5.6 Management review
Top management responsibility
documented planned intervals
Evaluate the need of change to the QMS
  Risk based approach
NOTE: written procedure and records required

## Slide 40: 5.6.2 Review input
a) feedback;
  From customers
  From predicate devices
b) complaint handling;
c) reporting to regulatory authorities;
d) audits;
e) monitoring and measurement of processes;
f) monitoring and measurement of product;

## Slide 41: 5.6.2 Review input
g) corrective action; NOT IN 9001:2015
h) preventive action
i) follow-up actions from previous management reviews;
j) changes that could affect the quality management system;
k) recommendations for improvement;
l) applicable new or revised regulatory requirements.

## Slide 42: Inputs from other clauses
Status of work environments (clause 6.4.1)
Information from servicing activities (clause 7.5.4)
DHF Change control (clause 7.3.10)
Other chosen by organization

## Slide 43: 5.6.3 Review output
Improvement
  Of system
  Of product
Changes for new regulatory requirements
Resources needed

## Slide 44: Documented procedure for HR (unvaried) 6.2 Human resources
establishing competence,
providing needed training,
  achieve or maintain the necessary competence
  Check effectiveness (risk based approach)
ensuring awareness of personnel
NOTE: records of education, training, skills and experience

## Slide 45: Written procedure 6.3 Infrastructure
Requirements for the infrastructure
  a) buildings, workspace and associated utilities;
  b) process equipment (both hardware and software);
  c) supporting services
Ensure that infrastructures are adequate for processes, prevent product mixup
Define maintenance
NOTE: Keep records of maintenance

## Slide 46: 6.4.1 Work environment 6.4.2 Contamination control
Requirements for work environment
Requirements for health, cleanliness and clothing of personnel
Processes to maintain adequacy
Reference norms ISO 14644 (cleanrooms) and ISO 14698 (biocontamination control)
Prevent contamination of work environment or other product

## Slide 47: 7.1 Planning of product realization
document one or more processes for risk management in product realization NEW
Plan product realization
quality objectives/ product specs
Equipment, infrastructure and work environment;
verification, validation, monitoring, measurement, inspection and test, product acceptance
handling, storage, distribution and traceability
records

## Slide 48: Process risk management
Process criticality level
Required product verification/ quality control
Required process validation
Training
Equipment maintenance
Work environment maintenance

## Slide 49: 7.2.1 Determination of requirements related to product
Specified by customer
  Delivery
  Post delivery
Not specified by customer, defined by the intended use
Regulatory
User training

## Slide 50: 7.2.2 Review of requirements related to product
Product ID and requirements
  Regulatory requirements
  traceability
Contract and order
User training NEW
Change control
NOTE: records of review

## Slide 51: 7.2.3 Communication
To customers
  a) product information;
  b) enquiries, contracts or order handling, including amendments;
  c) customer feedback, including complaints;
  d) advisory notices.
To regulatory authorities NEW
  Meddev

## Slide 52: 7.3 Design and development
Documented procedure
  Plan
  Input
  Output
  Review
  Verification
  Validation
  Transfer NEW
  Change control
Documented DHF to be kept updated with changes (clause 7.3.10) NEW

## Slide 53: 7.3.2 Design and development planning
a) the design and development stages;
b) the review(s) for each stage;
c) the verification, validation, and design transfer activities at each stage;
d) the responsibilities and authorities
e) traceability of outputs to inputs;
f) the resources (including competence)

## Slide 54: 7.3.3 Design and development inputs
defined per class
  a) functional, performance, usability (IEC 62366), safety
  b) regulatory requirements and standards;
  c) applicable output(s) of risk management;
  d) information derived from previous designs;
  e) other requirements
reviewed for adequacy and approved.
  complete, unambiguous, able to be verified or validated, and not in conflict with each other.

## Slide 55: 7.3.4 Design and development outputs
Meet inputs
Provide appropriate info
  Adequate for subsequent verification and validation
Specify essential product specs
  List of materials
  Shape/ design
  Packaging
Refer to QC criteria

## Slide 56: 7.3.5 Design and development review
Systematic review
  At predefined stages
Evaluate ability of results to match inputs
Propose corrective actions
representatives of functions concerned with the design and development stage being reviewed, as well as other specialist personnel

## Slide 57: 7.3.6 Design and development verification
Aim: ensure that the outputs have met the inputs
Document verification plans that include methods, acceptance criteria and, as appropriate, statistical techniques with rationale for sample size NEW
Verification of connections as applicable

## Slide 58: 7.3.7 Design and development validation
Ensure that device meets the intended use
  methods, acceptance criteria, and, as appropriate, statistical techniques with rationale for sample size NEW
  Representative product NEW
Clinical evaluation/ performance evaluation as integral part of validation
Completed prior to product release

## Slide 59: 7.3.8 Design and development transfer NEW
Transfer of design and development outputs to commercial manufacturing
  Issue of manufacturing work instructions
  Process capability
Process validation
QC processes validation

## Slide 60: 7.3.9 Control of design and development changes
Evaluation of impact of change
  function, performance, usability,
  Safety
  applicable regulatory requirements
  Existing product
Review, verification, validaiton
Formal approval and release

## Slide 61: 7.4.1 Purchasing process
Selection (risk based) NEW
  Supplier performance
  Product impact
Monitoring and continuous evaluation
NC control proportionate NEW to supplier risk class
NOTE: records of evaluation, selection, monitoring and re-evaluation

## Slide 62: 7.4.2 Purchasing information
Information on product
  Specs
  Quality levels for acceptance
Information on supplier
  QMS
  Personnel training
As applicable, written agreement that the supplier notify the organization of changes NEW
NOTE: records for traceability

## Slide 63: 7.4.3 Verification of purchased product
Extent of verification as per risk assessment NEW
  Supplier performance
  Product
  Changes of purchased products and their impact
Verification at supplier shall be described in the quality contracts
NOTE: records of verification and change control

## Slide 64: 7.5.1 Control of production and service provision (unvaried)
Production control (manufacturing, labeling, packaging, release, post delivery activites)
Qualification of infrastructure
Monitoring and control
  Monitoring and control equipment
Traceability as per Serial Number or Batch number

## Slide 65: 7.5.2 Cleanliness of product (unvaried) 7.5.5 Particular requirements for sterile medical devices
Control of cleanliness status
  Prior or alternate to sterilization
Control of cleaning processes
Removal of process residues
Can be manufactured out of cleanroom prior of (validated) cleaning procedures
Traceability of sterilization batch

## Slide 66: 7.5.3 Installation activities 7.5.4 Servicing activities
Performance and validation of installation
Control of external providers of installation/ servicing processes
Analysis of servicing records for NC control and improvement

## Slide 67: 7.5.6 Validation of processes for production and service provision
For high risk processes
  resulting output cannot be or is not verified
  deficiencies become apparent only after the product is in use or the service has been delivered
Demonstrate the ability of these processes to achieve planned results consistently
Required for sterilization and for packaging (clause 7.5.7)

## Slide 68: Validation SOP
criteria for review and approval of the processes;
equipment qualification
qualification of personnel;
methods, procedures and acceptance criteria;
as appropriate, statistical techniques (rationale for sample sizes) NEW
revalidation, including criteria for revalidation;
approval of changes to the processes NEW
NOTE: records required

## Slide 69: SW validation (see also clause 4.1.6)
SW used in production and service provision
Risk based approach NEW to validation of SW used in production
  At first use
  Change control

## Slide 70: 7.5.8 Identification (unvaried)
Throughout the whole device life
  production, storage, installation and servicing
Identification of QC status
  Includes cleanliness status
UDI number as per applicable regulation
Identification of returned product

## Slide 71: 7.5.9 Traceability (unvaried)
As per regulatory requirements
Detailed requirement for implantables
  records of components, materials, and conditions for the work environment
  Traceability by distributors
Detailed requirements for customer property identification and state of compliance

## Slide 72: 7.5.11 Preservation of product
during processing, storage, handling, and distribution
protect product from alteration, contamination or damage
  packaging design (and validation) NEW
  Labeling with storage instructions
  NOTE: record of special storage conditions

## Slide 73: Equipment SOP (unvaried) 7.6 Control of monitoring and measuring equipment
Monitoring and measurement consistent with product quality criteria
Equipment
  Identified
  Calibrated and recalibrated
  Protected by tampering
  Protected during storage and handling
Control of instrument NC

## Slide 74: 8 Measurement, analysis and improvement
Product conformity
QMS conformity
QMS continuous effectiveness
Appropriate (statistical) methods
Appropriate depth of controls

## Slide 75: Feedback SOP NEW? 8.2.1 Feedback
Feedback on whether the organization
  Has met customer requirements
  Post production information as per regulatory requirements
Feedback output is the input of risk management
  Product conformity
  Improvement

## Slide 76: Complaint SOP 8.2.2 Complaint handling
a) receiving and recording information;
b) evaluating information, determine if is a complaint;
  Justification of not acceptance of complaint
c) investigating complaints;
  Including impact of external parties
d) determining the need to CA reporting
e) handling of complaint-related (NC product) product;
f) determining the need to initiate corrections or corrective actions

## Slide 77: 8.2.3 Reporting to regulatory authorities NEW
the organization shall document procedures for providing notification to the appropriate regulatory authorities.
NOTE: appropriate records

## Slide 78: 8.2.4 Internal audit
Compliance to
  planned and documented arrangements,
  requirements of this International Standard,
  quality management system requirements established by the organization,
  applicable regulatory requirements
    Also product requirements
Effective implementation and maintenance

## Slide 79: Audit SOP (unvaried)
Responsibilities (impartiality and competence)
  Planning
  conducting audits
  recording and reporting
Audit plan as per risk approach
  Process criticality
  Results from previous audits
Reporting of NC found during audits
Follow up of CA

## Slide 80: 8.2.5 Monitoring and measurement
Of processes (clause 8.2.5)
  Suitable methods for monitoring
    Production records and evaluation of collected data
Of products (clause 8.2.6)
  Evidence of conformity with the acceptance criteria
  Identify of control personnel
    For release
    For any control for implantables
  Quarantine of products under QC evaluation

## Slide 81: NC product SOP 8.3 Control of nonconforming product
Aim: NC product is identified and controlled to prevent its unintended use or delivery.
Identification and documentation
Segregation,
Evaluation,
  determination of the need for an investigation
  notification of any external party responsible for the NC
Disposition of nonconforming product
  Rationale of decisions

## Slide 82: 8.3.2 Actions in response to NC product detected before delivery
NC correction (reword see clause 8.3.4)
Limitation to use
Use under concession
  Justified
  Approved by customer
  Still in compliance with regulatory requirements
  Detailed record (of justification and responsible board)

## Slide 83: Regrade for use under concession
3.12.4 regrade: alteration of the grade of a NC product in order to make it conform to requirements differing from the initial requirements
3.6.3 grade: category or rank given to different requirements for an object having the same functional use

## Slide 84: 8.3.3 Actions in response to NC product detected after delivery
action appropriate to the effects, or potential effects
document procedures for issuing advisory notices
  At any time
NOTE: new clause but same content as old version

## Slide 85: SOP for rework NEW 8.3.4 Rework
takes into account the potential adverse effect of the rework
procedures shall undergo the same review and approval as the original procedure
product shall be verified NEW
  Increased sampling?

## Slide 86: Some definitions ISO 9000:2015
3.12.8 rework: action on a NC product to make it conform to the requirements
  Rework can affect or change parts of the nonconforming product or service.
3.12.9 repair: action on a NC product to make it acceptable for the intended use
  does not necessarily make the product or service conform to the requirements, a concession may be required.
  Includes repair during maintenance.
  can affect or change parts of the NC product

## Slide 87: “Management review input” SOP 8.4 Analysis of data
determine, collect and analyse appropriate data
  a) feedback;
  b) conformity to product requirements;
  c) characteristics and trends of processes and product including opportunities for improvement;
  d) suppliers;
  e) audits;
  f) service reports, as appropriate

## Slide 88: 8.5 Improvement
Ensure consistent
  suitability, adequacy and effectiveness of the quality management system
  Medical device safety and performance
Inputs for improvement
  quality policy, quality objectives,
  audit results,
  postmarket surveillance
  analysis of data, corrective actions, preventive actions and management review

## Slide 89: 8.5.2 Corrective action
a) reviewing nonconformities (including complaints);
b) determining the causes of nonconformities;
c) evaluating the need for action to ensure that NCs do not recur;
d) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation;
e) verifying that the corrective action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; NEW
f) reviewing the effectiveness of corrective action taken

## Slide 90: 8.5.3 Preventive action
a) determining potential NCs and their causes;
b) evaluating the need for action to prevent occurrence
c) planning and documenting action needed and implementing such action, including, as appropriate, updating documentation;
d) verifying that the action does not adversely affect the ability to meet applicable regulatory requirements or the safety and performance of the medical device; NEW
e) reviewing the effectiveness of the preventive action taken, as appropriate.